## Slide 1: Perché superare i PMT a dinodi
FLUTTUAZIONI NEL GUADAGNO DEL PRIMO DINODO:rende impossibile photon counting.
BASSA CAPACITA’ DI DISTINGUERE IL SINGOLO FOTONE
LINEARITA’ LEGATA AL GUADAGNO
TT ELEVATO
TTS ELEVATO
STRUTTURA COMPLESSA E COSTOSA
ALTA SENSIBILITA’ AI CAMPI MAGNETICI
ALTO CONSUMO
PESSIMO FUNZIONAMENTO A TEMPERATURE CRIOGENICHE
ALTA RADIOATTIVITA’ METALLO DIFFUSO (dinodi)

## Slide 2
Rivelatori di Luce (UV-IR)
luce
Amplificazione
luce
corrente
Amplifi
cazione
Assente/
debole
Alte
Amplificazioni
Bassi livelli
di luce
a
vuoto
No/basse
Amplificazioni
Alti livelli
di luce
a
stato
solido
nasce nel 1913 ed e’ commerciale RCA nel 1936
Tecnologia in forte sviluppo
corrente
106-106

## Slide 4
Affinché i fotoni vengano assorbiti e producano coppie elettrone-lacuna, la loro energia deve essere maggiore della bandgap del materiale, cioè deve essere hν>Eg, o, equivalentemente, hc/λ>Ec.
Per il Silicio, Eg=1.12eV e quindi risulterebbe λg=1.11μm.
coefficiente di assorbimento del materiale
Il suo inverso, 1/α, è detto lunghezza di penetrazione ed è la distanza entro la quale circa il 63% dei fotoni viene assorbito.
Nella costruzione o nella scelta di un fotodiodo, deve essere dunque scelto un materiale la cui lunghezza di penetrazione alla lunghezza d’onda da rivelare sia di dimensioni dell’ordine della regione di svuotamento.

## Slide 5
Photodetectors
Absorption Coefficient and Photodiode Materials
Absorbed Photon create Electron-Hole Pair.
Cut-off wavelength vs. Energy bandgap
Absorption coefficient
Incident photons become absorbed as they travel in the semiconductor and light intensity decays exponentially with distance into the semiconductor.


## Slide 6: Absorption Coefficient
Absorption coefficient α is a material property.
Most of the photon absorption (63%) occurs over a distance 1/α (it is called penetration depth δ)

## Slide 7
Spectral Responsitivity (Sk)
Quantum efficiency (Q.E.) ρ

## Slide 8
Per i fotodiodi si definisce efficienza quantica il rapporto tra il numero di coppie prodotte e raccolte agli elettrodi e il numero di fotoni incidenti:
Si definisce, inoltre, spectral responsitivity del fotodiodo il rapporto tra la fotocorrente prodotta e il potere ottico incidente:
Ideal Photodiode

## Slide 9
Relative high Z useful for photon detection via photoelectric effect

## Slide 10
1,12 eV
intrinsic semiconductor: no good radiation detector-thermic current-low sensitivity

## Slide 13
≈1022 atoms/cm3

## Slide 21
In a pure intrinsic (undoped) material the electron density n and hole density p are equal. n = p = ni
In this volume
we have 4.5 108 free charge carriers, but only 3.2 104 e-h pairs produced by a M.I.P.
 Reduce number of free charge carriers,
i.e. deplete the detector (drogaggio, giunzione e polarizzazione inversa)
Come ottenere un segnale?

## Slide 23
1,12 eV
Questa extra lacuna va ad occupare una posizione all’interno della gap proibita, spostando la gap più verso la banda di valenza.
Typical doping:
1013 atoms/cm3 on a density of
1022 atoms/cm3

## Slide 26
C

## Slide 27
p+
p+
p+
n

## Slide 28
Heavily doped semiconductor as electrical contacts: 1020 atoms/cm3 highly conductive
p+,n+ sign

## Slide 30
+
+
-
-

## Slide 31
Difficulty to achieve uniform multiplication
in entrance windows.
Gain strong function of applied voltage and
temperature.
In drift region no multiplication.
If all incident radiation interacts within drift
region: all electrons are subject to the same
multiplication in high field region.
Statistical variation in avalanche process
introduce fluctuations in multiplication
lowering the energy resolution.
+
-

## Slide 34: Impact Ionization
p+
n+
-V
E
V < VAPD => photodiode 1 collected pair/generated pair
VAPD < V <VBD => APD <M> collected pairs/generated pair
V > VBD => Geiger-mode APD inf. collected pairs/generated pair
n
Simulated diode reverse current
IMPACT IONIZATION ON
IMPACT IONIZATION OFF
VBD
VAPD
diode structure
electric field
in the reversed
bias diode
depletion width

## Slide 35
La ricerca di nuovi fotorivelatori porta inevitabilmente
ai dispositivi al silicio
p-i-n
Amplificazione
1
ed oggi + ?
Diodi p-n
APD
Amplificazione
50-100
luce
luce
A questo punto rispetto ai
classici PMT mancano:
-superficie attiva
-guadagno per bassi livelli di luce

## Slide 36
matrici di microcelle in parallelo
ogni microcella: GM-APD + Rquenching
V. M. Golovin and A. Sadygov
MEPHI
Giunzioni pn in regime Geiger G-APD
-High gain as PMT
-small active surface

## Slide 37
First a thin layer (2 to 4 mm) is grown epitaxial on a wafer (typically 300 mm thick)
The first step towards a p-n junction is the diffusion of dopants with high concentration into the epitaxial layer p+,n+.
The junction is then formed some 0.5 mm below the surface by a shallow diffusion (or ion implantation) with the opposite dopant.
In a final step an extremely thin but heavily doped layer is created on the surface which distributes the potential uniformly over the whole area of the diode.
Passivation of the surface with SiO2.
Techniques of photolithography are used to remove selectively areas oxide where entrance window are to be located
-
-
+
+

## Slide 40
Al
ARC
-Vbias
Back contact
p
n+
p
n+

Rquenching
h
p+ silicon wafer
Front contact
- Vbias
n pixels
One pixel
fired
Two pixels fired
Three pixels
fired
Current (a.u.)
Time (a.u.)
Out
SiPMT / G-APD / MPPC
microcontatori indipendenti in regime geiger limitato
Geiger spento dalla caduta del campo e resistenza di
quenching R.
Segnale in uscita somma segnali singoli geiger
V Geiger 10-20% sopra la tensione di brekdown 25-60 V
Rquench
Altezza impulso

## Slide 43
SiPM: da Hamamatsu MPPC
100 euro!!

## Slide 45
Quenching
resistor 1x1
1600 pixel
=145 Ω
microcel
=230 kΩ
Quenching
resistor 3x3
14400 pixel
=26 Ω
microcel
=380 kΩ

## Slide 46
Caratteristiche I – V, polarizzazione inversa dei fotorivelatori più comuni, dipendenza del guadagno fotoconduttivo G dalla polarizzazione
Fotorivelatori:
PIN
APD
SiPM

## Slide 47
Metal-Resistor-Semiconductor
R superficiale disaccoppiamento diodi
30-80 MΩ cm.
RSi microcell quenching ~ 300KΩ
C microcell ~ 50 fC
RC ~ durata segnale ~ 15-20 ns
R
R
- pixel
SiPM: Tecnica costruttiva
ITC-irst
+
+
Notes: Aggiungere guadagno IRST, crosstalk,buchi nelle celle

## Slide 48
E Si-PM = e geom (1-R)(1-eax)e GM
= QE (g) / dE/dx (mip)
geom: area geometrica sensibile. Aree morte necessarie per resistori
    di isolamento per ridurre cross talk. Ampio sviluppo tecnologico ~0,6
QE: assorbimento nella regione di svuotamento: per l=400nm (3,10 eV)
a=5,4x106 cm-1, 99,9% in 2mm. Ottimizzazioni per g e mip.
spessore di svuotamento, compromesso fra assorbimento e corrente
oscura (meglio per mip <1m si riduce corrente oscura).
R: coefficiente di riflessione (5%)
sviluppo tecnologico
e GM : efficienza trigger Geiger (90-100%)
SiPM: efficienza

## Slide 49
+
-

## Slide 51: Geiger-mode APD
Diode biased at VD > VBD
p+
n-
-VD
n
t < t0 i=0 (if no free carriers
in the depletion region)
t = t0 carrier initiates the avalanche
t0 < t < t1 avalanche spreading
t > t1 self-sustaining current
In order to be able to detect another photon, quenching mechanism needed:
t
i
t0
i=iMAX
t1
VBIAS
VD
Two solutions:
large resistance:
passive quenching
analog circuit:
active quenching
[extended literature from politecnico
di Milano (Cova et al.)]
t
i
t0
t1
t2
t
vD
vBD

## Slide 52: Passive quenching (1)
Studied in detail in the ’60 to model micro-plasma instabilities
[ R.J. McIntrye, JAP vol.32, n.6 1961; R. Haitz, JAP vol.35, n.5 1964 ]
OFF condition:
switch open, capacitance charged to VBIAS,
no current flowing
P01 = turn-on probability
probability that a carrier traversing the
high-field region triggers the avalanche
ON condition
The GM-APD can be modeled with an electrical circuit and
two probabilities.
CD
RS
VBD
RQ
VBIAS
VD
DIODE
RQ = quenching resistance
RS = series resistance (emulates
the avalanche spreading)
CD = diode capacitance
VBD = breakdown voltage
VBIAS = bias voltage

## Slide 53: Passive Quenching (2)
CD
RS
VBD
RQ
VBIAS
IINT
ID
VD
DIODE
ON condition:
avalanche triggered, switch closed
the capacitance discharges to VBD with
a time constant RSxCD, at the same time
the diode current grows to (VBIAS-VBD)/RQ
P10 = turn-off probability
probability that the number of carriers
traversing the high-field region
fluctuates to zero.
OFF condition:
switch open
diode capacitance recharghes from VBD
to VBIAS with a time constant RQxCD
Ready for new detection

## Slide 55: GAIN
The area of the current pulse represents the gain
GAIN = Area/q = IMAX x t / q = C x (VBIAS-VBD)/q
The capacitance should be large in order to have
a high gain.
On the other hand, a large capacitance leads to longer
recovery times.
VBIAS-VBD
GAIN
C2>C1
C1

## Slide 59
+
-

## Slide 60
+
-
+

## Slide 61
The noise rate decreases as the temperature becomes lower. The temperature coefficient of noise rate at 0.5 p.e. threshold is −5 %/◦ C. There is a factor 2 reduction of the dark count every 8°C

## Slide 66
cell1
cell2
cell1
cell2
Solutions:
operate at low over-voltage => low gain => few photons emitted
optical isolation structure:

## Slide 68
Afterpulses

## Slide 70
Effetto del crosstalk ottico e degli after – pulses sull’andamento temporale della fototensione di uscita di un SiPM e sulla sua risoluzione temporale
S = risposta elettrica del SiPM ad un impulso luminoso costituito da un singolo fotone
d = risposta elettrica del SiPM ad un impulso luminoso costituito da un singolo fotone: durante questa fotorivelazione si ha l’attivazione spuria di una seconda microcella, limitrofa a quella colpita dal fotone; a causa dell’interferenza ottica fra i 2 pixels la risposta del SiPM sovrastima l’intensità ottica incidente

## Slide 73
Esercizi

## Slide 74
Example
(b) What is the thickness of a Ge and In0.53Ga0.47As crystal layer that is needed for absorbing 90% of the incident radiation at 1.5 mm?
For Ge,   5.2  105 m-1 at 1.5 m incident radiation.
For In0.53Ga0.47As,   7.5  105 m-1 at 1.5 m incident radiation.
For Ge,   5.2  105 m-1 at 1.5 m incident radiation.
For In0.53Ga0.47As,   7.5  105 m-1 at 1.5 m incident radiation.
(b)

## Slide 75
Example
0
0.2
0.4
0.6
0.8
1
800
1000
1200
1400
1600
1800
Wavelength (nm)
Responsivity (A/W)
The responsivity of an InGaAs pin photodiode
InGaAs pin Photodiodes
Consider a commercial InGaAs pin photodiode whose responsivity is shown in fig. Its dark current is 5 nA.
What optical power at a wavelength of 1.55 mm would give a photocurrent that is twice the dark current? What is the QE of the photodetector at 1.55 mm?
What would be the photocurrent if the incident power in a was at 1.3 mm? What is the QE at 1.3 mm operation?

## Slide 76
Solution
(a) At l = 1.55´10-6 m, from the responsivity vs. wavelength curve we have R  0.87 A/W. From the definition of responsivity,
we have
From the definitions of quantum efficiency h and responsivity,

Note the following dimensional identities: A = C s-1 and W = J s-1 so that A W-1 = C J-1. Thus, responsivity in terms of photocurrent per unit incident optical power is also charge collected per unit incident energy.

## Slide 77
Solution
(b) At l = 1.310-6 m, from the responsivity vs. wavelength curve, R = 0.82 A/W.
Since Po is the same and 11.5 nW as in (a),
The QE at l = 1.3 mm is

## Slide 78
Example
Bandgap and photodetection
(a) Determine the maximum value of the energy gap which a semiconductor, used as a photoconductor, can have if it is to be sensitive to yellow light (600 nm).
(b) A photodetector whose area is 510-2 cm2 is irradiated with yellow light whose intensity is 20 mW cm-2. Assuming that each photon generates one electron-hole pair, calculate the number of pairs generated per second.
Solution
(a) Given, l = 600 nm, we need Eph = hu = Eg so that,
Eg = hc/l = (6.62610-34 J s)(3108 m s-1)/(60010-9 m) = 2.07 eV
(b) Area = 510-2 cm2 and Ilight = 2010-3 W/cm2.
The received power is
P = Area Ilight = (510-2 cm2)(2010-3 W/cm2) = 10-3 W
Nph = number of photons arriving per second = P/Eph
= (10-3 W)/(2.0591.6021810-19 J/eV)
= 2.97871015 photons s-1 = 2.97871015 EHP s-1.

## Slide 79
(c) For GaAs, Eg = 1.42 eV and the corresponding wavelength is
l = hc/ Eg = (6.62610-34 J s)(3108 m s-1)/(1.42 eV1.610-19 J/eV)
= 873 nm (invisible IR)
The wavelength of emitted radiation due to EHP recombination is 873 nm.
(d) For Si, Eg = 1.1 eV and the corresponding cut-off wavelength is,
lg = hc/ Eg = (6.62610-34 J s)(3108 m s-1)/(1.1 eV1.610-19 J/eV)
= 1120 nm
Since the 873 nm wavelength is shorter than the cut-off wavelength of 1120 nm, the Si photodetector can detect the 873 nm radiation (Put differently, the photon energy corresponding to 873 nm, 1.42 eV, is larger than the Eg, 1.1 eV, of Si which mean that the Si photodetector can indeed detect the 873 nm radiation)
Example
Bandgap and Photodetection
(c) From the known energy gap of the semiconductor GaAs (Eg = 1.42 eV), calculate the primary wavelength of photons emitted from this crystal as a result of electron-hole recombination. Is this wavelength in the visible?
(d) Will a silicon photodetector be sensitive to the radiation from a GaAs laser? Why?
Solution