## Slide 1
Optical Module in NEMO Phase-2
A glass sphere 13” (Vitrovex):
Single large area photomultiplier : Hamamatsu 10” PMT R7081
Optical gel : Waker SilGel 612
μ-metal wire cage
PMT base circuit : ISEG PHQ7081-i-2m modified
FEM (Front End Module) electronic board
System for timing calibration (TIM-CAL)
13” OM sketch: lateral view
TIM-CAL
FEM
ISEG

## Slide 2
Assembled OM
Picture of an assembled hemisphere: glass, PMT, Gel, mu-metal cage, ISEG

## Slide 3
Glass sphere
standard 13 inches deep-sea instrumentation vessels in borosilicate glass, produced by Vitrovex
two half spheres: ½ transparent, ½ painted black
no vacuum valve
unique penetration for the 12-pin connector
Refractive index | 1.48 (>350 nm)
Transmission | >95% (>350 nm)
Density at 20 0C | 2.23 g cm-3
Thermal conductivity | 1.2 W m-1K-1
Characteristics of 13 inches spheres | 
Depth rating (m) | 10000
Overall diameter (mm) | 330
Wall thickness (mm) | 11
Mass (kg) | 7.89
Buoyancy (empty) (N) | 114
Diameter shrinkage per 1000m depth (mm) | 0.30
Active base
+5 Volts supply (bipolar voltage supply before modification)
Cathode-1^dynode and 1^dynode-anode voltages individually controllable
Anode current max : 100 microAmpere
Power consumption : 150mW @ 2000 Volts

## Slide 4
Mechanical support for FEM electronic board and TIM-CAL
The mounted FEM
The mounted TIM-CAL

## Slide 5
A cage of mu-metal wire was chosen as magnetic shield (ITEP, Moscow):
a hemispherical part ( 30 cm diameter, 14 cm height)
a flat part (30 cm diameter ) with a hole in its centre ( 12 cm diam.)
wire of 1 mm of diameter
pitch of 68 x 68 mm
shadow on the photocathode ≈ 5%
average shielding factor measured ≈ 4
Magnetic shield
Picture of the parts of the cage
The cage around the 10” PMT

## Slide 6
Soldering of the ISEG base on the PMT
Mechanical Frame
Base positioning
Base soldering
Wires cut off
The end
OM assembly procedure: base soldering

## Slide 7
Pre-pulses appear in the 10–80 ns interval time before the main pulse and take the place of the main pulse. They arise from a direct photo-effect on the first dynode due to the photons which pass the photocathode without photoelectron conversion.
Late pulses appear in the 10–80 ns interval time after the main pulse and take the place of main pulse. They are caused by backscattered photoelectrons from the first dynode.
Type 1 after pulses appear in the 10–80 ns interval time after the main pulse, in the wake of the main pulse. They are produced by a luminous reaction on the electrodes bombarded by the electrons.
Type 2 after pulses arrive in the 80 ns–16ms interval time after the main pulse, in the wake of the main pulse. They are caused by the ionization of residual gases left or formed inside the PMT. These gas atoms can be ionized by electrons and, since they are of opposite charge, will accelerate back towards the photocathode where they can produce further photoelectrons. Common ions are H2+ , He+ , CH4+ , and Cs+.
Spurious pulse
G=5.107

## Slide 9
Afterpulses

## Slide 10
Measurements of the effects of the Earth’s magnetic field
Every PMTs was measured in 3 inclinations:
- vertical downwards ( Tilt = 0° ) ; 50° downwards ( Tilt = 50° ); horizontal ( Tilt = 90° )
NORTH
PMT downwards
Dy 1
Dy 2
90°
270°
180°
0°
Starting position for each PMT
For each inclination, the PMT under test was rotated 360° around its vertical axis in 30° steps
Each PMT started its rotation from the same position with respect to the box and to the Earth’s magnetic field
North
Z
90°
270°
180°
Horizontal position
Tilt : 90°
Vertical axes
0°
Horizontal angles
Tilt
North
0°
Horizontal
angles
90°
270°
180°
Z
Vertical position
Tilt : 0°
Vertical
axes
North
0°
Horizontal
angles
90°
270°
180°
Tilt : 50°
Z
Vertical
axes
Top view

## Slide 11
Detection efficiency
In naked 8” PMTs the impact of the magnetic field was smaller than naked 10”
The shield reduced considerably the variations for the 10” PMT
The increased QE in the HQE 8” PMT compensates the smaller detection area respect to the 10”
0°
50°
90°
NAKED SHIELDED
0°
50°
90°

## Slide 12
Gain
Greatest gain variation was less than 10% for both the naked 8”
Considerable variations in the 10” PMT naked, up to 29%
The shield reduces variations in both the 8” PMTs, with variations less than 4.4 %
The shield reduces strongly the variations in the 10” PMT, with variation less than 7%
NAKED SHIELDED

## Slide 13
Peak to Valley ratio
Considerable variations for all the un-shielded PMTs
Great reductions in variation by using the magnetic shield
Small improvements in average values with the magnetic shield
NAKED SHIELDED

## Slide 14
Charge Resolution (sigma)
Large effects of magnetic field for the unshielded 10” PMT
The mu-metal cage greatly reduced the variation in 10” PMT
No large effects due to the magnetic field for 8” PMTs.
The mu-metal cage reduced variations for both 8” PMTs
NAKED SHIELDED

## Slide 15
TT
No significant variations due to magnetic field for all the PMTs
No considerable effects of the mu-metal cage
Difference in average value due to the different size between 8” and 10” PMT
NAKED SHIELDED

## Slide 16
TTS
Variations over the 10% in all the un-shielded PMTs
Considerable reduction in variations with the magnetic shield
No significant improvement in average values with magnetic shield
NAKED SHIELDED

## Slide 17
The segmentation of the detection area in the DOM will aid in distinguishing single-photon from multi-photon hits. Moreover, two-photon hits can be unambiguously recognised if the two photons are registered in separate tubes.

## Slide 18
XP2600
QUASAR-370
MOTIVATION
The segmentation of the detection area in the DOM will aid in distinguishing single-photon from multi-photon hits. Moreover, two-photon hits can be unambiguously recognised if the two photons are registered in separate tubes.

## Slide 19
Quasar-370 phototube has excellent time and
very good single electron resolutions
no prepulses
no late pulses in TTS
low level of afterpulses
~100% effective collection effiency
1 ns TTS (FWHM)
very good SER (competitive to HPD)
immunity to terrestrial magnetic field
>2 sensitivity

## Slide 20: Quasar-370LSO with LSO crystal
1996-97, ICRC1997
Single electron resolution ~35%
(fwhm)
Jitter ~ 1 ns (fwhm)
The PMT used in the Quasar-370LSO had ~17% (eff)

## Slide 23: The Digital Optical Module (DOM)
Multi-PMT DOM
31 small PMTs
Almost uniform coverage
Photon counting
Minimize pressure transitions
All electronics inside

## Slide 24
Multi-PMT Optical Module for KM3
In Mediterranean sea
31 3” PMTs inside a 17” glass sphere
Cooling shield and stem
First prototype end of 2011
Single vs multi photon hit separation
Larger photocade area per OM
PMT solution: very coarse method to do photon
counting, charge proportional to time over
threshold. Poor time resolution
VSiPMT solution: true photon counting. Good
time resolution
Photon counting useful to:
-reduce background: isotropy of radioactivity
potassium.
-event: directionality
Cerenkov
directional light
Fondo b da K40
Isotropic Cerenkov light
-Photon counting
-Time resolution

## Slide 26
75 mm PMT
4 manufacturers
  ETEL ( 25 delivered )
  Hamamatsu ( 10 delivered )
  MELZ ( first prototype expected )
  Zhan Chuang Photonics
High QE (>32%@380nm;>22%@470nm)
Extension to ~90 mm diameter with concentrator ring
ETEL

## Slide 28
Expantion cone

## Slide 30: Ultima versione PMT

## Slide 31: R12199 characteristics

## Slide 33
New design HV with <35 mW power consumption
12 PMTs in top
19 PMTs in bottom
Front matched to sphere
Supported by foam cores via concentrator ring
Optically coupled with optical gel

## Slide 34
Orientation of PMT’s. The
positive axis points upward

## Slide 36
Power provided by two copper conductor
All data transmission and slow control communication requires a single optical fibre

## Slide 39
Front End
Instrumentation
Power management
Data Management

## Slide 41: DOM